TENDER SPECIFICATIONS (for main items)

1 HPPP Unit:

Bench-top unit.

Anodized aluminum frame and panels made of painted steel.

The unit includes wheels to facilitate its mobility.

Main metallic elements made of stainless steel.

Diagram in the front panel with distribution of the elements similar to the real one.

Includes:

Pelton turbine:

Speed range: 0 – 1600 rpm.

Nominal power: 1 kW.

Torque: 6 Nm.

Number of buckets: 20.

2 x servo-driven injectors.

Generator / Brake:

Output power: 1 kW.

Nominal voltage: 3PH, 400VAC + N + GND.

Speed range: 0 – 1600 rpm.

Pulley diameter: 290 mm.

Effective radius: 160 mm.

Computer controlled water pump:

Maximum pressure: 8 bar.

Maximum water flow: 106 l/min – 5 bar.

Nominal power: 3 kW 3PH.

Frequency drive:

Nominal power: 3 kW.

Pressure sensor:

SP-1: 0 – 30 psi (0 – 2 bar).

SP-2, SP-3, SP-4: 0 – 100 psi (0 – 6.7 bar).

Load cell - force sensor: 0 – 50 N.

Flow sensor: 5 – 150 l/min.

Speed sensor: 0 – 10000 rpm.

Water transparent tank:

Capacity: 130 l approx.

Manual drain valve.

Main manual valve.

Pressure gauge: 0 – 15 bar.

Two x servo-driven valve.

Security pressure switch.

The complete unit includes as well:

Advanced Real-Time SCADA.

Open Control + Multicontrol + Real-Time Control.

Specialized EDIBON Control Software based on LabVIEW.

National Instruments Data Acquisition board (250 KS/s, kilo samples per second).

Calibration exercises, which are included, teach the user how to calibrate a sensor and the importance of checking the accuracy of the sensors before taking measurements.

Projector and/or electronic whiteboard compatibility allows the unit to be explained and demonstrated to an entire class at one time.

Capable of doing applied research, real industrial simulation, training courses, etc.

Remote operation and control by the user and remote control for EDIBON technical support, are always included.

Totally safe, utilizing 4 safety systems (Mechanical, Electrical, Electronic & Software).

Designed and manufactured under several quality standards.

Optional ICAI software to create, edit and carry out practical exercises, tests, exams, calculations, etc. Apart from monitoring user's knowledge and progress reached.

This unit has been designed for future expansion and integration. A common expansion is the EDIBON Scada-Net (ESN) System which enables multiple students to simultaneously operate many units in a network.

Additional recommended elements (Not included):

- HPPP-CR. Control and Regulation Unit for HPPP.

Three phase power supply: 400 VAC + 3PH.

Nominal power: 2 kW.

ON / OFF switch.

Emergency stop push button.

Reset alarm push button.

Operation signals state:

Ready for operation indicator.

Start request indicator.

Sync. Conditions indicator.

Permission for sync. Indicator.

Measurement terminals.

Automatic Speed and Voltage Controller: Turbine controller.

Analog instrumentation panel:

Generator voltmeter.

Measurement range: 0 – 500 VAC.

Generator wattmeter.

Measurement range: 0 – 1,5 kW.

Generator varmeter.

Measurement range: 0-1,5 kVAr.

Fuses: 3 x 10 A, 1 x 2 A.

Power strip: 5 x 400 VAC.

Network analyzer:

Nominal voltage: 400 VAC.

Network analyzer display. It shows:

Active, reactive and apparent power.

Active, reactive and apparent energies.

Lines and phase currents.

Line and phase voltages.

Frequencies.

Power factor.

Automatic servo injector controller:

This device, known as “Automatic Turbine and Generator Controller”, automates the opening and closing process of the two turbine injectors, as well as the synchronization of the generator with the electrical grid. Thanks to its operation, a completely autonomous control of the electrical power generated by the Pelton turbine is achieved.

In addition to its main function, this controller acts as a multifunction protection, monitoring several critical parameters of the electrical generator. Among these parameters are generator speed, voltage, frequency, power and phase balance. Continuous monitoring of these parameters ensures safe and efficient operation of the generator, protecting it from abnormal operating conditions that could damage the unit or affect the quality of the power supply.

In short, this device is essential for the comprehensive control and protection of the Pelton turbine and associated electrical generator, ensuring its optimal operation and providing reliable and stable electrical power production.

Protection system:

Generator:

Maximum/minimum voltage: 59/27.

Maximum/minimum frequency: 81O/U.

Voltage asymmetry.

Detection of dead busbars.

Overvoltage: 32.

Load unbalance: 46.

Negative sequence power/reduced power: 32R/F.

Overcurrent by defined curve: 50/51.

Inverse time overcurrent: IEC255.

Measured ground fault: 50N/51N.

Phases rotation.

Switches faults.

Network:

Maximum/minimum voltage (59/27).

Maximum/minimum frequency (81O/U).

Vector jump.

Phase rotation.

- N-REVT/1K. 1kW Three-Phase Variable Resistors Module.

Parallel, series, star and delta connection.

Resistance: 3 x 750 Ohm.

Current: 3 x 2 A.

Nominal power: 1000 W.

- N-CAR19T/3C. 3x300 var Three-Phase Configurable Capacitors Module.

Three three-phase banks with capacitors of 2 µF each one.

Configurable star and delta connection.

Nominal voltage: 400 VAC.

Nominal power: 3 x (3 x 300) Var.

2 HPPP/CIB. Control Interface Box:

The Control Interface Box is part of the SCADA system.

Control interface box with process diagram in the front panel.

The unit control elements are permanently computer controlled.

Simultaneous visualization in the computer of all parameters involved in the process.

Calibration of all sensors involved in the process.

Real time curves representation about system responses.

All the actuators’ values can be changed at any time from the keyboard allowing the analysis about curves and responses of the whole process.

Shield and filtered signals to avoid external interferences.

Real time computer control with flexibility of modifications from the computer keyboard of the parameters, at any moment during the process.

Real time computer control for parameters involved in the process simultaneously.

Open control allowing modifications, at any moment and in real time, of parameters involved in the process simultaneously.

Three safety levels, one mechanical in the unit, another electronic in the control interface and the third one in the control software.

3 DAB. Data Acquisition Board:

The Data Acquisition board is part of the SCADA system.

PCI Express Data acquisition board (National Instruments) to be placed in a computer slot.

Analog input: Channels= 16 single-ended or 8 differential. Resolution=16 bits, 1 in 65536. Sampling rate up to: 250 KS/s (kilo samples per second).

Analog output: Channels=2. Resolution=16 bits, 1 in 65536.

Digital Input/Output: Channels=24 inputs/outputs.

The Data Acquisition board model may change at any moment, providing the same or better features than those required for the unit.

4 HPPP/CCSOF. Computer Control +Data Acquisition+Data Management Software:

The three softwares are part of the SCADA system.

Compatible with the industry standards.

Flexible, open and multicontrol software, developed with actual windows graphic systems, acting simultaneously on all process parameters.

Management, processing, comparison and storage of data.

Sampling velocity up to 250 KS/s (kilo samples per second).

Calibration system for the sensors involved in the process.

It allows the registration of the alarms state and the graphic representation in real time.

Open software, allowing the teacher to modify texts, instructions. Teacher’s and student’s passwords to facilitate the teacher’s control on the student, and allowing the access to different work levels.

This unit allows the 30 students of the classroom to visualize simultaneously all the results and the manipulation of the unit, during the process, by using a projector or an electronic whiteboard.

5 Cables and Accessories, for normal operation.

6 Manuals:

This unit is supplied with 8 manuals: Required services, Assembly and Installation, Interface and Control software, Starting-up, Safety, Maintenance, Calibration & Practices manuals.

*References 1 to 6 are the main items: HPPP + HPPP/CIB + DAB + HPPP/CCSOF + Cables and Accessories + Manuals are included in the minimum supply for enabling normal and full operation.

EXERCISES AND PRACTICAL POSSIBILITIES TO BE DONE WITH THE MAIN ITEMS

1.-  Determination of pelton turbine operating parameters.

2.-  Analysis of the relationship between injection pressure and turbine mechanical power.

3.-  Determination of Pelton turbine efficiency as a function of water jump, injector opening and load applied to the turbine.

4.-  Analysis of the influence of water flow on turbine performance: the water flow rate supplied to the turbine is adjusted and its performance is measured in relation to the mechanical power applied to the turbine shaft.

5.-  Study of the influence of water injection asymmetry in the pelton turbine on its performance.

6.-  Study the relationship between injector opening and generated mechanical power: measuring of the mechanical power developed by the turbine while varying the injector openings.

7.-  Optimization of turbine performance: testing with different combinations of water head height and injector opening to find the configuration that maximizes turbine performance.

8.-  Study the behavior of sensors: Analyze the response of flow rate, pressure, velocity, and torque sensors under different operating conditions.

Additional practical possibilities with additional recommended elements HPPP-CR, N-REVT/1K and N-CAR19T/3C:

9.-  Study of energy production in stand-alone mode.

10.-Study of the impact of the connection to the generator of electrical loads in island mode. Influence on generator voltage and frequency.

11.-Study of the synchronization process of the generator with the grid. Control of frequency, voltage and phase sequence.

12.-Study of the causes and consequences of the motorization of the generator in synchronism with the grid.

13.-Study of the impact of sudden uncoupling of the turbine and the grid on the electrical and mechanical variables of the turbine.

14.-Analysis of the effect of load on power plant performance: Study of the synchronization of the electrical generator with the grid. Use of the network analyzer to measure and visualize the most relevant electrical parameters of the process.

15.-Study of the synchronization of the electric generator with the grid. Use of the network analyzer to measure and visualize the most relevant electrical parameters of the process.

16.- Evaluation of system protection and safety: verification of advanced controller protection functions, such as over-frequency, over-voltage, over-current and reverse power protection. Verification of the system response to fault or overload situations.

17.-Analysis of system dynamics using SCADA: Use SCADA to visualize electrical and mechanical parameters in real time. Monitor in detail how the parameters vary and the interactions between them during different plant operations.

Other possibilities to be done with this unit:

18.-Many students view results simultaneously.

To view all results in real time in the classroom by means of a projector or an electronic whiteboard.

19.-Open Control, Multicontrol and Real Time Control.

This unit allows intrinsically and/or extrinsically to change the span, gains; proportional, integral, derivative parameters; etc, in real time.

20.-The Computer Control System with SCADA allows a real industrial simulation.

21.-This unit is totally safe as uses mechanical, electrical/electronic, and software safety devices.

22.-This unit can be used for doing applied research.

23.-This unit can be used for giving training courses to Industries even to other Technical Education Institutions.

24.-Control of the HPPP unit process through the control interface box without the computer.

25.-Visualization of all the sensors values used in the HPPP unit process.

- Several other exercises can be done and designed by the user.

TENDER SPECIFICATIONS (for optional items)

a) Technical and Vocational Education configuration

7 HPPP/ICAI. Interactive Computer Aided Instruction Software.

This complete software package consists of an Instructor Software (EDIBON Classroom Manager - ECM-SOF) totally integrated with the Student Software (EDIBON Student Labsoft - ESL-SOF). Both are interconnected so that the teacher knows at any moment what is the theoretical and practical knowledge of the students.

- ECM-SOF. EDIBON Classroom Manager (Instructor Software).

ECM-SOF is the application that allows the Instructor to register students, manage and assign tasks for workgroups, create own content to carry out Practical Exercises, choose one of the evaluation methods to check the Student knowledge and monitor the progression related to the planned tasks for individual students, workgroups, units, etc...so the teacher can know in real time the level of understanding of any student in the classroom.

Innovative features:

• User Data Base Management.

• Administration and assignment of Workgroup, Task and Training sessions.

• Creation and Integration of Practical Exercises and Multimedia Resources.

• Custom Design of Evaluation Methods.

• Creation and assignment of Formulas & Equations.

• Equation System Solver Engine.

• Updatable Contents.

• Report generation, User Progression Monitoring and Statistics.

- ESL-SOF. EDIBON Student Labsoft (Student Software).

ESL-SOF is the application addressed to the Students that helps them to understand theoretical concepts by means of practical exercises and to prove their knowledge and progression by performing tests and calculations in addition to Multimedia Resources. Default planned tasks and an Open workgroup are provided by EDIBON to allow the students start working from the first session. Reports and statistics are available to know their progression at any time, as well as explanations for every exercise to reinforce the theoretically acquired technical knowledge.

Innovative features:

• Student Log-In & Self-Registration.

• Existing Tasks checking & Monitoring.

• Default contents & scheduled tasks available to be used from the first session.

• Practical Exercises accomplishment by following the Manual provided by EDIBON.

• Evaluation Methods to prove your knowledge and progression.

• Test self-correction.

• Calculations computing and plotting.

• Equation System Solver Engine.

• User Monitoring Learning & Printable Reports.

• Multimedia-Supported auxiliary resources.

8 HPPP/FSS. Faults Simulation System.

Faults Simulation System (FSS) is a Software package that simulates several faults in any EDIBON Computer Controlled Unit.

The "FAULTS" mode consists in causing several faults in the unit normal operation. The student must find them and solve them.

There are several kinds of faults that can be grouped in the following sections:

Faults affecting the sensors measurement:

- An incorrect calibration is applied to them.

- Non-linearity.

Faults affecting the actuators:

- Actuators channels interchange at any time during the program execution.

- Response reduction of an actuator.

Faults in the controls execution:

- Inversion of the performance in ON/OFF controls.

- Reduction or increase of the calculated total response.

- The action of some controls is annulled.

On/off faults:

- Several on/off faults can be included.

b) Multipost Expansions options

9 MINI ESN. EDIBON Mini Scada-Net System for being used with EDIBON Teaching Units.

MINI ESN. EDIBON Mini Scada-Net System allows up to 30 students to work with a Teaching Unit in any laboratory, simultaneously.

The MINI ESN system consists of the adaptation of any EDIBON Computer Controlled Unit with SCADA integrated in a local network.

This system allows to view/control the unit remotely, from any computer integrated in the local net (in the classroom), through the main computer connected to the unit.

Main characteristics:

- It allows up to 30 students to work simultaneously with the EDIBON Computer Controlled Unit with SCADA, connected in a local net.

- Open Control + Multicontrol + Real Time Control + Multi Student Post.

- Instructor controls and explains to all students at the same time.

- Any user/student can work doing "real time" control/multicontrol and visualisation.

- Instructor can see in the computer what any user/student is doing in the unit.

- Continuous communication between the instructor and all the users/students connected.

Main advantages:

- It allows an easier and quicker understanding.

- This system allows you can save time and cost.

- Future expansions with more EDIBON Units.

The system basically will consist of:

This system is used with a Computer Controlled Unit.

- Instructor’s computer.

- Students’ computers.

- Local Network.

- Unit-Control Interface adaptation.

- Unit Software adaptation.

- Webcam.

- MINI ESN Software to control the whole system.

- Cables and accessories required for a normal operation.