TENDER SPECIFICATIONS

Bench mounted spray drier for processing aqueous emulsions, solutions, suspensions and colloids. This unit is suitable for aqueous solutions only.

Diagram in the front panel with a distribution of the elements similar to the real unit.

Downward co-current operation (a fine jet of the product is brought into contact with a hot air stream).

The unit components are made of glass with gasket free ground glass flanges.

The unit includes a panel at the top of the frame with 4 displays for the temperature sensors, switch on/off for the feed pump and feed pump speed controller, switch on/off for the fan and switch on/off for the heating element (resistance).

Characteristics:

Maximum drying capacity: 1000 ml/h approx.

Temperature range: 40 - 200ºC (temperature at inlet).

Dry air volume range: 0.2 - 0.65 m³/min.

Spray air pressure range: 0.5 - 2.5 kg/cm².

Feed pump volume range: 102 - 1800 ml/h approx.

Maximum air pressure: 70 mbar.

The chemically resistant powder coated housing includes the fan and heating element (resistance).

All clamps and fittings are designed to allow assembly and removal of the glass elements rapidly and easily.

Fan:

Power: 0.4 kW.

Velocity: 3000 rpm.

Drying air throughflow: 70 m³/h (fixed).

Heating element (resistance) of 3 kW.

Drying chamber:

Material: borosilicate glass.

It includes a spray nozzle, diameter: 0.5 mm.

The spray assembly incorporates a manual de-blocking device that prevents the jet nozzle from becoming blocked.

Feed pump: peristaltic pump.

Cyclone:

Material: borosilicate glass.

Sample collection bottle:

Material: hard glass.

Volume: 500 ml.

Exhaust tube:

Outer diameter: 50 mm.

Filter/air regulator located between the compressor (not included) and the unit to ensure that the drying air does not include contaminants.

Four temperature sensor type "J" to measure:

Environmental temperature.

Air inlet temperature in the drying chamber.

Exhaust air temperature.

Feed temperature.

Glass vessel, volume: 1  l.

Cables and Accessories, for normal operation.

Manuals: This unit is supplied with following manuals: Required Services, Assembly and Installation, Starting-up, Safety, Maintenance & Practices Manuals.

- Dimensions: 500 x 500 x 1200 mm approx. (19.68 x 19.68 x 47.24 inches approx.) - Weight: 80 Kg approx. (176.36 pounds approx.).

EXERCISES AND PRACTICAL POSSIBILITIES

1.- Operation principle of a spray drier.

2.- Effect of the drop size on the drying process.

3.- Effect of the air input temperature on the drying process.

4.- Effect of the feed flow of the product on the drying process.

5.- Mass balance of a spray drier.

6.- Spray drier efficiency.

Additional practical possibilities:

7.- Energy balance of a spray drier.

Optional

SSPB/ICAI. Interactive Computer Aided Instruction Software:

This complete software package consists of an Instructor Software (EDIBON Classroom Manager - ECM-SOF) totally integrated with the Student Software (EDIBON Student Labsoft - ESL-SOF). Both are interconnected so that the teacher knows at any moment what is the theoretical and practical knowledge of the students.

- ECM-SOF. EDIBON Classroom Manager (Instructor Software).

ECM-SOF is the application that allows the Instructor to register students, manage and assign tasks for workgroups, create own content to carry out Practical Exercises, choose one of the evaluation methods to check the Student knowledge and monitor the progression related to the planned tasks for individual students, workgroups, units, etc... so the teacher can know in real time the level of understanding of any student in the classroom.

Innovative features:

• User Data Base Management.

• Administration and assignment of Workgroup, Task and Training sessions.

• Creation and Integration of Practical Exercises and Multimedia Resources.

• Custom Design of Evaluation Methods.

• Creation and assignment of Formulas & Equations.

• Equation System Solver Engine.

• Updatable Contents.

• Report generation, User Progression Monitoring and Statistics.

- ESL-SOF. EDIBON Student Labsoft (Student Software).

ESL-SOF is the application addressed to the Students that helps them to understand theoretical concepts by means of practical exercises and to prove their knowledge and progression by performing tests and calculations in addition to Multimedia Resources. Default planned tasks and an Open workgroup are provided by EDIBON to allow the students start working from the first session. Reports and statistics are available to know their progression at any time, as well as explanations for every exercise to reinforce the theoretically acquired  technical knowledge.

Innovative features:

• Student Log-In & Self-Registration.

• Existing Tasks checking & Monitoring.

• Default contents & scheduled tasks available to be used from the first session.

• Practical Exercises accomplishment by following the Manual provided by EDIBON.

• Evaluation Methods to prove your knowledge and progression.

• Test self-correction.

• Calculations computing and plotting.

• Equation System Solver Engine.

• User Monitoring Learning & Printable Reports.

• Multimedia-Supported auxiliary resources.